## Slide 1
FWPD Webinar
Intelligent Pulse Detection
Byron Gao
APAC(China) Application Engineer

## Slide 2: Agenda
Intelligent pulse detection
  Pulse Detection Solution & Supported Products
  Pulse Detection Application Examples
  FWPD Pulse Detection Feature Overview
    Digital Baseline Stabilization (DBS)
    Detection Window
    Coincidence
    Intelligent Pulse Detection
    Hysteresis
    Real-time or Custom Pulse Analysis
    Data collection and Data Multiplexer
    Timestamp
    Latency Control

## Slide 3: Pulse Detection Solution
High performance digitizers ADQ7 and ADQ14
  Standard FW FWDAQ included
Pulse Detection FW FWPD for ADQ7 and ADQ14
  FW FWPD is an add on option
Pulse Characterization GUI application
Example application code in C and Python
FWPD User’s Guides for ADQ7 and ADQ14
Pulse Detection application note

## Slide 4: FWPD – On ADQ7 and ADQ14
USB3.0
MTCA.4
PXIe
PCIe
10GbE
SSPCIe
Flexible choice of product and form factor

## Slide 5: Pulse Detection Application Examples
Big Physics
Time-of-flight
Scientific instruments
LIDAR
RADAR
Neutron time-of-flight

## Slide 6: Pulse Detection Overview
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Create a stable and smooth baseline
Channel trigger is conditioned by activity on other channels
Real-time analysis in the FPGA
Ensures a constant minimum data rate from the digitizer
Ignore pulses outside the detection window
Dynamic recording, dynamic level trigger, and zero suppression
Flexible detection and presentation of metadata
Timestamp for real-time timing

## Slide 7: Digital Baseline Stabilization (DBS)
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
DBS tracks slow and periodic baseline variations originating from:
Temperature variations
Component aging
Pattern noise from ADC due to interleaving
DBS helps avoid false readings and/or performance degradation:
Missed pulses
Erroneous detection of false pulses
Incorrect analysis of pulse characteristics

## Slide 8: Digital Baseline Stabilization with DBS
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Baseline variation over time

## Slide 9: Pattern Noise Reduction with DBS
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Multiple interleaved ADC cores with different offset
Results in a periodic zig-zag pattern in the output:
0 10 | 20 | 30 | 40
3,5
3
2,5
2
1,5
1
0,5
0
-0,5
-1
-1,5
ADC
ADC
ADC
ADC
Interleaved digital output

## Slide 10: Pattern Noise Reduction with DBS, continued
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Significant pattern noise suppression with DBS
Pattern noise with DBS
Pattern noise without DBS

## Slide 11: Detection Window
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Events outside the detection window ignored

## Slide 12: Example of use - Detection Window
Use window to synchronize to a shot of particles
Select reasonable flight-time only
Discriminate outliers
Source
Target
Detector

## Slide 13: Coincidence
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Coincidence used for channel-dependent triggering
Trigger event on green channel only if it occurs within a configurable timeslot (the coincidence window) after the red channel is triggered.

## Slide 14: Example of use - Coincidence
Study flight time of particles
Multiple sets of detectors
Particles/pulses only stored if hitting multiple detectors within a certain time frame
Other pulses are discarded to reduce data to be transferred/stored
Timestamp is used to determine flight time between detectors

## Slide 15: Example of use - Coincidence, continued
Time-of-flight example

## Slide 16: Intelligent Pulse Detection vs Standard Data Acquisition
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Individual channel and dynamic data-driven record length
  Fixed vs variable record lengths
Fixed record length
Variable record length

## Slide 17: Intelligent Pulse Detection
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
The pulse start is determined by the trigger level.
The pulse end is determined by the reset level.
Trigger and reset levels relative to stable baseline
  Determines beginning and end of pulse data record

## Slide 18: Example of use – Dynamic Data-Driven Record Length
Any radiation source has a random generation of particles
Optimize recording/transfer to only include meaningful data
Example application - Thomson scattering in fusion reactors

## Slide 19: Valid Trigger/Reset Points
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Signal may contain noise that results in false trigger/reset
Safety mechanism required that prevents false events

## Slide 20: Trigger/Reset Arm Hysteresis
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Help with trade-off between accepting false events and missing out on real pulses

## Slide 21: Trigger/Reset Arm Hysteresis, cont’d
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Help with trade-off between accepting false events and missing out on real pulses

## Slide 22: Intelligent Pulse Detection – Edges
DBS
Detection window
Coincidence
Pulse analysis
Data collection
Latency control Timestamp
Leading Edge Window (LEW) before trigger point Trailing Edge Window (TEW) after reset point
LEW
TEW
Trigger level
Reset level
Pulse detection
Region of interest
Mechanism to enable study of pulse edges

## Slide 23: Examples of use – LEW/TEW
Pulse length vary on a radiation decay
If pulses (particles) pile up the length of the pulse vary
Optimal transmission of data

## Slide 24: Real-time Pulse Analysis
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Time of peak
Width
LEW
TEW
Peak
Built-in real-time pulse analysis
  Peak value, time over threshold (width), peak timestamp

## Slide 25: Examples of use – Pulse Analysis
Peak detection in time-of-flight mass spectrometry (TOFMS)
Area is energy

## Slide 26: Custom Pulse Analysis
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
LEW
TEW
Open FPGA enables custom real-time pulse analysis
  Examples – area, power, custom peak definition, Gaussian curve fit, spline interpolation, qualify/disqualify pulse, etc.
Available via firmware development kit or design service

## Slide 27: Examples of end use – Custom Pulse Analysis
Custom processing in real-time optimized for any application
See application notes for examples and further information

## Slide 28: Data Collection
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Five data collection modes for different purposes:
Overview summary
Mode | Trigger Types | Detection Window | Data Type | Data Padding | Nth Record
0 | All | No | Raw | No | No
1 | Level | Yes | Meta | Yes | No
2 | All | No | Raw | No | Yes
3 | Level | Yes | Raw | Yes | No
4 | Level | No | Raw | Yes | No

## Slide 29: Data Collection
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Five data collection modes for different purposes:
0. Default. Most features turned off.
Fixed length raw data records with any trigger source
Variable length raw data records with level trigger
1. Pulse metadata within detection window is collected.
Packets of metadata are structured into records
Same as Mode 0, but only every Nth record is collected.
Similar to Mode 1, but with raw data instead of metadata.
  Variable length records of raw pulse data
Same as Mode 3, but without detection window.

## Slide 30: Data Collection, cont’d
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Collected data is stored into records
Both data and header records created
Timestamp included in header
Records are transferred to host
Some modes can be combined
Multiplexer with meta & raw data channels
For example mode 1 & 3

## Slide 31: Intelligent Data Collection with Data Multiplexer
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
FPGA
Raw data A
ADC
AFE
Standard pulse char.
Pulse detection
Custom pulse char.
ADC
AFE
Standard pulse char.
Custom pulse char.
Pulse detection
Pulse characteristics A
Record frame
Record frame
MUX
3
1
A
B
PC
Use cases:
Normal operation
Raw data A Pulse char. A
3
1

## Slide 32: Examples of use – Data Multiplexer
This is typically used during application development
Compare raw data with analyzed for development/debugging

## Slide 33: Intelligent Data Collection with Data Multiplexer
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
FPGA
Raw data A
ADC
AFE
Standard pulse char.
Pulse detection
Custom pulse char.
ADC
AFE
Standard pulse char.
Custom pulse char.
Pulse detection
Raw data B
Record frame
Record frame
MUX
3
3
A
B
PC
Use cases:
Normal operation
Raw data A Raw data B
3
3

## Slide 34: Intelligent Data Collection with Data Multiplexer
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
FPGA
Use cases:
Normal operation
Pulse char. A Pulse char. B
Pulse characteristics A
ADC
AFE
Standard pulse char.
Pulse detection
Custom pulse char.
ADC
AFE
Standard pulse char.
Custom pulse char.
Pulse detection
Pulse characteristics B
Record frame
Record frame
MUX
1
1
1
1
A
B
PC

## Slide 35: Data Collection - Mode 3 and 4 (raw data)
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Data
Header
Discarded
Discarded
Discarded
Raw data of pulses 1 and 2 are collected in two separate records. All other data is discarded to optimize the amount of data processed and transferred to the PC.
LEW
LEW
TEW
TEW
Pulse #1 Pulse #2
Record

## Slide 36: Data Collection - Mode 3 and 4 with LEW/TEW Overlap
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Record #1
Record #2
LEW
TEW
TEW
LEW
LEW
TEW
Trigger 1
Trigger 2
If LEW and TEW overlap, the record is extended to contain both pulses in the same record

## Slide 37: Data Collection - Mode 1 (metadata)
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Peak 16b
Width 16b
Time of peak 32b
Time-stamp 64b
Peak 16b
Width 16b
Time of peak 32b
Pulse characteristics are stored in metadata packets
Header Pulse #1 Pulse #2 #3 #n
Time-stamp 64b
Cust. value1 x bits
Cust. value2 y bits
…
Header Pulse #1 Pulse #2
Cust. value1 x bits
Cust. value2 y bits
…
#3 #n
Metadata record
Header
…

## Slide 38: Timestamp
Header timestamp if detection window
Header timestamp if no detection window
Peak timestamp definition dependent on Detection Window
Peak
Detection window
Peak timestamp if no detection window
Peak timestamp if detection window
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Trigger

## Slide 39: Timestamp – Header without Detection Window
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
LEW
TEW
TEW
LEW
Pulse #1
Timestamp in header
Timestamp in header
Raw data #1
Raw data #2
Header
Header
Header
Header
Raw data channel
Pulse char. channel
Pulse char. #1
Pulse char. #2
Pulse #2

## Slide 40: Timestamp – Header with Detection Window
DBS
Detection window
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Timestamp
Coincidence
External trigger
Timestamp reset
Pulse #1
Pulse #2
Timestamp in header
Timestamp in header
LEW
TEW
TEW
LEW

## Slide 41: Controlling Pulse Data Latency
In data-driven acquisition the latency is high if pulses are sparse
This is sometimes not acceptable
One example is updating plots/diagrams in a graphical user interface
A time-out mechanism required to force data transfer if time between consecutive pulses is too long
Padding helps fill the data pipeline/FIFO with dummy data (zeros)
Padding is recognized as dummy data by the application software and discarded

## Slide 42: Latency Control - Padding with Detection Window
DBS
Detection window
Coincidence
Pulse detection
Pulse analysis
Data collection
Latency control Timestamp
Padding helps ensure constant data rate

## Slide 43: Intelligent Pulse Detection Demo
Demo video to review some FWPD features discussed
FWPD feature overview video also available here

## Slide 44: Application example: Pulse Characterization GUI

## Slide 45: Further Documentation Information
FWPD datasheet for ADQ7 here and ADQ14 here
FWPD user guide for ADQ7 here and ADQ14 here
FWPD website here
“The art of pulse detection” application note, here

## Slide 46: Contacts
Byron Gao
APAC(China) Application Engineer
Peng.Gao@Teledyne.com